## Weekly QC
Weekly MRI Equipment Quality Control For Large Phantom
MRI Facility Name |  |  |  |  | MRI Scanner Identifier | [9 empty] | column | criteria | criteria
1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 |  |  | 2 | -5 | 5
Date | Setup and Table Position Accuracy |  | Center Freq (Hz) | TX Gain or Attenuation (db) | Geometric Accuracy Measurements (Axial Slice #5 Diameter) |  |  | High Contrast Spatial Resolution (Slice 1-highest resolved) |  | Low-Contrast Detectability (Slice 8) | Artifact Evaluation | Inspections/ Clean Equipment | Tested By | Notes | 6 | 146 | 150
| Accuracy Ok? | Console Ok? |  |  | H/F Sagittal Localizer Length (mm) | A/P (mm) | R/L (mm) | Upper Left | Lower Right | # of Spokes | Any Present? | Bore & Sponges Cleaned, Coils Inspected |  |  | 7 & 8 | 188 | 192
Action Limits | ± 5mm | Yes/No | ± 256 | 120 ± 20% | 148 ± 2 mm | 190 ± 2mm | 190 ± 2 mm | ≤ 1.0mm | ≤ 1.0mm | 7 | Yes/No | Yes/No | Initals |  | 9 | 1
[15 empty] | 11 | 7
[15 empty] | Yes | No
[15 empty] | Yes | Yes
[15 empty] | 5 | 96 | 144 | based on +/-20% with average of 120
Qualified Medical Physicist | [5 empty] | Date

## Visual Checklist
Weekly MRI Accreditation Program Visual Checklist
Facility |  |  |  |  | MRI Scanner Identifier
| Patient Transport and Gantry |  |  |  | Filming Viewing |  | RF Integrity and Control Room | [5 empty] | Facility Safety |  |  |  | MRI Room |  |  | Initials
Date | Table Position and Other Displays | Alignment Lights | Horizontal Smoothness of Motion and Stability | Vertical Motion Smoothness and Stability | Laser Camera | Light Boxes | RF Door Contacts | RF Window-Screen Integrity | Operator Console Switches/Lights/Meters | Patient Monitor (If Present) | Patient Intercom | Room Temperature/Room Humidity | Emergency Cart | Safety Warning Signage | Door Indicator Switch (If Installed) | Cryogen Level Indicator | Bore Cleaned | Sponges Clean and No Tears | No Fraying Coils |  | KEY
[21 empty] | P | PASS
[21 empty] | F | FAIL
[21 empty] | N/A | N/A
Qualified Medical Physicist | [5 empty] | Date